## Slide 1: Results based on 1D and 2D arrays at IPHC Strasbourg
Nico Brosda, Stéphane Higueret 19.11.2024

## Slide 2: Schedule (I‘ll leave room for discussion after each block)
Short Reminder: How we measure at Cyrcé
Reminder: First arrays measured at Cyrcé (partially shown at Lacassagne meeting)
Comparison: Difference between bonding fanout on GaN and direct bonding
Results: Voltage of AMS readout circuit + Linearity checks
Results: Homogeneity of array response
Some details of the image generation
Results: Imaging with different arrays
Results: Imaging in time dependence (creating movies of the beam)
Results: Signal in dependence of the proton energy
Results: Comparison with Gafchromic scans
Outlook: What results I haven‘t shown / What‘s to come

## Slide 3: Reminder Setup at Cyrcé
Protons with ≈ 24 MeV
Usual current ≈ 2 nA
GaN arrays on XY-translation stage
Optional: Diffuser in proton beam line (homogenize beam)
Diodes read out by AMS Test Kit
We do not vary the beam any longer (as done at the first measurements with 1 s pulses)
We usually integrate for 3 s
Objects to image
Aftwerwards: Data processing and image generation (code written by me)

## Slide 4: Results until June this year
Excellent Linearity
We can resolve the
energy of protons
(and thus biological objects)
Ability to image objects (or the undiffused beam)

## Slide 5: Difference in array design afterwards
Previous 64 diode arrays
Maximum total length 500µmx140=7cm
Direct bonding between
PCB card and diode
SINGLE DIODE
500µm
200µm
500µm
400µm
400µm
30µm
10µm
10µm
20µm
“Fan out” from diode contacts on the GaN wafer PCB card and diode
All measurements following are done with arrays without the ”fan out” = fewer fabrication steps + direct bonding!

## Slide 6: Comparison between array designs
“Old” 1x64 0.5x0.5 mm2 array:
“New” 1x128 0.5x0.5 mm2 diodes:

## Slide 7: Let’s just compare the raw data
“Old” 1x64 0.5x0.5 mm2 array:
“New” 1x128 0.5x0.5 mm2 diodes:

## Slide 8: Why the difference?
Better quality (fewer fabrication steps)
Impact of metal ”fan out” on GaN wafer
We observe a transfer of irradiation-induced electrons between lines of “fan out”!
(We also modelled and partially corrected for this)
Before
After correction
Due to the new design, we do not have this problem any longer

## Slide 9: Overview of arrays measured so far
With “fan out”:
3 different 1x64 0.5x0.5 mm2 line arrays
Without ”fan out”:
1x128 0.5x0.5 mm2 line array
1x128 0.5x0.25 mm2 line array
2x64 0.5x0.5 mm2 line array (0.25 mm shifted)
2 different 11x11 1x1 mm2 matrix arrays
11x11 0.5x0.5 mm2 matrix arrays

## Slide 10: Reminder: Sample structure
Active region

## Slide 11: 50 Hz noise on undiffused beam?!
Compared is the highest and lowest (dark) signal of arbitrary measurements with and without diffuser
I observed this modulation for different measurement days whenever we measured without diffuser; I will nevertheless bring these observations in plots for you to compare
If you count the periods on the right side you find: 10 periods in 200 ms meaning 50 in 1s = 50 Hz…

## Slide 12: Part II
Voltage tests
Linearity checks
Homogeneity of array response

## Slide 13: Voltage tests

## Slide 14: Sensitivity and signal quality

## Slide 15: All voltage scans up to the 22nd of October

## Slide 16: Sensitivity and signal quality
Note that the sensitivity implies that we could e.g. detect a current ~ 200 lower than we had for the big matrix arrays)
Meaning: 2 nA at Faraday cup / 200 = 10 pA at Faraday cup under these conditions for the large diodes.

## Slide 17
Ratio between the signals
other measurement day ≠ comparable…
SINGLE DIODE
500µm
200µm
500µm
400µm
400µm
30µm
10µm
10µm
20µm
500µm
250µm
200µm
400µm
170µm
SINGLE DIODE

## Slide 18
Ratio between the signals
SINGLE DIODE
500µm
200µm
500µm
400µm
400µm
30µm
10µm
10µm
20µm
70µm
800µm
200µm
1mm

## Slide 19: What will be analyzed
We measured the linearity voltage-dependent for currents between 0.5 nA – 2.5 nA (5 points)
This is not the biggest range, but it is enough for some quick statements about the linearity of arrays (especially when collecting measurements as a benchmark for new arrays)
Again only diodes with signal will be considered! Signal = netto signal (signal –dark)
I will plot the mean signal of all diodes (black, x) and the signal of each diode (grey, x)
Also plotted: Dark noise (black, v); mean of signal std (blue, ^); std of diodes signal (green, ^)
I also show the inaccuracy of the mean signal = Sqrt[mean of signal std2 + std of diodes‘ signal2] (black, ^)
I fit a linear function of type f(x) = a*x to each diode's response and collect their R2 value – in the plot I will show the mean of these fits (red, -- line) and the mean of all R2 values.
I also fit a function of type f(x) = a*x1/2 + (dark noise) to each signal noise, again I will show the mean (red, -- line) and the mean of all R2 values.

## Slide 20: Linearity in dependence of voltage
Small matrix array

## Slide 21
Small matrix array

## Slide 22: Comparison with voltage
Small matrix array
Big matrix array #1

## Slide 23: Homogeneity in diode response
Every diode is measured at several (overlapping) positions in the beam
1x128 0.5x0.5 mm2 array

## Slide 24: Factor for normalizaton
1x128 0.5x0.5 mm2 array – 2 measurements at different times
We are in ≈ 6% diode response homogeneity

## Slide 25: Comparison with old array design
Old design had ≈ 120 % deviation, both 128 line arrays are within ≈ 6%
Only one not working diode out of 2 times 128 diodes ≈ efficiency of 99,6 %

## Slide 26: Same for 2x64 array / matrix arrays
Not every diode is overlapping at every position!
2x64 0.5x0.5 mm2 array

## Slide 27: Factor for each line
2x64 0.5x0.5 mm2 array
Note that there is a pattern underneath both curves
The proton current decreased during the measurement
This low-frequency pattern will be filtered in the future!
There seems to be a difference between the two lines?!
After aligning lines to global mean

## Slide 28: Homogeneity for matrix arrays
Small matrix array
Big matrix array #2

## Slide 29: Possible defect of readout circuit
In all of the latest measurements one channel 0 and one channel 2 x the normal signal… they`re neighbours

## Slide 30: Part III
Details on image generation
Imaging

## Slide 31: How I create the images
Out of translation

## Slide 32: How I create the images
Out of translation

## Slide 33: How I create the images
Out of translation + diode geometry

## Slide 34: Different techniques
I create our images (and in general all plots) with Python’s library Matplotlib:
Pcolourmesh = Different - sized rectangles with colour
Contourf = Values are assigned to coordinate points in between the colour is interpolated
Also possible (but no real gain in my tests): Imshow (equally sized pixels, standard for images)
I create our images (and in general all plots) with Python’s library Matplotlib:
Pcolourmesh = Different - sized rectangles with colour
Contourf = Values are assigned to coordinate points in between the colour is interpolated
Also possible (but no real gain in my tests): Imshow (equally sized pixels, standard for images)

## Slide 35: All colourmaps possible!

## Slide 36: General info about images I’ll show
Signals are for now shown in units of the AMS circuit output (unclear how to exactly calculate the current out of the AMS output amplitude)
Images are background subtracted and normalized (as shown)
In some cases, I might compensate for single lines not working
I stay with oversampling since I still need to implement a deconvolution into super-resolved pixels

## Slide 37: The round aperture / 200 µm diffuser
1x128 0.5x0.5 mm2 array
Real pixel information

## Slide 38: Sunrise of the proton beam 
1x128 0.5x0.5 mm2 array

## Slide 39: The three available diffusers
2x64 0.5x0.5 mm2 array
40 µm
200 µm
400 µm

## Slide 40: The misc shape
1x128 0.5x0.5 mm2 array
Real pixel information

## Slide 41
1x128 0.5x0.5 mm2 array
2x64 0.5x0.5 mm2 array

## Slide 42
1x128 0.5x0.5 mm2 array
The logo

## Slide 43
Smaller diodes 0.25 mm in scan direction + 0.25 mm steps
0.5 mm diodes + 0.5 mm steps
1x128 0.5x0.5 mm2 array
1x128 0.5x0.25 mm2 array

## Slide 44: Beam imaging without diffuser
1x128 0.5x0.5 mm2 array

## Slide 45
Quite some change of the beam within our measurement day!
Can we get a better time resolution? Live imaging?
Note that we cannot compare the signal here due to the different set voltages!
Smaller diodes 0.25 mm in scan direction + 0.25 mm steps
0.5 mm diodes + 0.5 mm steps

## Slide 46: Part IV
Imaging in time dependence = movies

## Slide 47: What was measured?
11x11 1x1 mm2 matrix #1 with higher time resolution (≈ 0.7 fps in this case; 1540 images over 36 minutes!)
Video will be x50 increased in speed
The array was fixed during the measurement.
Three phases:
    Beam stability (no changes)
    Changes with collimation magnets
    Wire moving through the beam
    (To get a beam profile Cyrcé facility moves wires through the beam line and measures a current. For us this means something shadows the array.)

## Slide 48
Three phases:
  Beam stability (no changes)
  Changes with collimation magnets
  Wire moving through the beam
11x11 1x1 mm2 matrix array #1

## Slide 49: What was measured?
11x11 1x1 mm2 matrix #2 with 1 fps
Video will be x30 increased in speed
The array was fixed during the measurement.
What limits faster measurements so far: The readout circuit, or more precisely the software we have to use in the test kit…

## Slide 50
Live diffuser in and out
Slit closing
Wire moving
Magnet shift
11x11 1x1 mm2 matrix array #2

## Slide 51
Square aperture
200 µm diffuser
Current increase from 100 pA to 3.6 nA
(Array is tilted a bit)
11x11 1x1 mm2 matrix array #2

## Slide 52: Part V
Resolving proton energy
PEEK wedge
Imaging of biological tissue

## Slide 53: The Bragg peak (1 map and y translated data)
1x128 0.5x0.5 mm2 array

## Slide 54: Wedge out of PEEK material

## Slide 55: Map of PEEK wedge
2x64 0.5x0.5 mm2 array

## Slide 56: What if we change the energy of the protons?
We have 20 wheel positions = 20 proton energies!

## Slide 57: Energy-dependent wedge measurement
1x128 0.5x0.5 mm2 array

## Slide 58: Biological tissue
I put this into this category because the mechanism to image tissue is based on our diodes’ energy resolution.
Disclaimer: No animals were harmed exclusively for our measurements!
Measurement with 2x64 array (1 no signal diode replaced, oversampling)

## Slide 60: Part VI
Comparison with Gafchromic Images

## Slide 61: Gafchromic images

## Slide 62: Method to obtain a comparable response

## Slide 63: Some qualitative comparisons
2x64 0.5x0.5 mm2 array

## Slide 64: Some qualitative comparisons
2x64 0.5x0.5 mm2 array

## Slide 65: High exposure?
2x64 0.5x0.5 mm2 array

## Slide 66: Energy resolution?
2x64 0.5x0.5 mm2 array

## Slide 67: And finally one quantitative comparison
1x128 0.5x0.5 mm2 array

## Slide 68
H
B
G
D
1x128 0.5x0.5 mm2 array

## Slide 69
H
B
G
D
Norm all to 1 and align x-axes to get a quantitative comparison!
Data from Gafchromic extracted with online tool – points aligned by hand
(without investing several hours this at least gives a nice idea of the data curve)
Looks excellent, and it seems that our normed response can be compared to the response of the gafchromic!
1x128 0.5x0.5 mm2 array

## Slide 70: Part VII
Outlook:
Results I haven‘t shown Future plans

## Slide 71: Not shown / results not ready yet
Energy dependent imaging
Voltage dependent imaging
Super-resolution: Can we improve our images by oversampling?
More movies / images
Gafchromic difference maps
How many samples do we need to get good measurements?
And even more different things

## Slide 72: Outlook
Using a simulation software for validation (GEANT4 interfaced by GATE)
    Deposited energy = signal?
    Projection on to other facilities (e.g. more energy)
    Contributions to energy deposition
    Possible damage channels
  Long-term irradiation damage studies
  Measuring at other facilities
  Detecting carbon ions?
  Important: Getting the new readout circuits!