## Slide 1: Latest Cyrcé results and ideas for the next measurements
Measured arrays:
128 x 0.5 x 0.5 mm2 array – ‘Array1‘
128 x 0.25 x 0.5 mm2 array - ‘Array3‘
Stéphane Higueret, Nico Brosda

## Slide 2: How we measured
round aperture of
proton beam
ø = (36,51 mm)
~ 30 mm in reality
Array ~ 128 x 0.5 mm = 64 mm
Mounted on xy - stage
Objects to image

## Slide 3: Details about the beam
Beam for all measurements (except beam imaging) diffused with 200 m aluminium diffuser (nearly flat beam)
Proton current ~ 2 nA on sample (meaning ~ 2.43 before diffuser)
But distortion due to additional collimator in beam line (when we expcected 2.00 nA on the sample, we had ~ 1.90 nA)
Energy of protons: 24.915 MeV (before diffuser)
24.007 MeV (behind diffuser)
23.776 MeV (at output into air)
Instability at start of measurement day (beam current drops of ~ 0.1 nA over 15 min)

## Slide 4: The 200 um diffuser used for these measurements!
H
B
G
D

## Slide 5: Samples measured
Active region

## Slide 6
Array 3
Array 2 (not measured yet)

## Slide 7: Arrays‘ dimensions
500µm
250µm
200µm
400µm
170µm
SINGLE DIODE
SINGLE DIODE
500µm
200µm
500µm
400µm
400µm
30µm
10µm
10µm
20µm
DIODES LINE (x140)
Maximum total length 500µmx140=7cm
128 diodes read out – here referred as: Array3
128 diodes read out – here referred as: Array1
DIODES LINE (x140)
Maximum total length 500µmx140=7cm
Most measurements were done with this array – if not explicitly mentioned the results are from Array1!

## Slide 8: IV-curves

## Slide 9: Voltage Scan
From 0 - 2 V for both arrays.
Comparison with old voltage scan
Signal = Mean from measurement
Noise = Std from measurement

## Slide 10: Comparison in one graph

## Slide 11: Ratios and netto signal

## Slide 12: The signal ratio between the 2 array?
We expect a ratio of
0.17 / 0.4 = 2,35
(The diodes have half the size, but the active area
is smaller than the diodes)
Only the 2 V seem the fit this expectation, sadly we cannot see how this trend continues.
There might be some conclusions concerning the effective size of the active area inside the sample!

## Slide 13: Images – Array 1

## Slide 14: Image raw vs background subtracted

## Slide 15: Fanout (old) vs direct bond (new)

## Slide 16: The difference in concept
Previous 64 diode arrays
Masks of new array generation – no fanout on GaN sample, direct bonding to PCB

## Slide 17: Round aperture

## Slide 18: Logo

## Slide 19: Misc shape

## Slide 20: Results y-translated measurements
What was measured?
How uniform are the diodes?
How good is the normalisation factor?
Comparison to factor from the previous measurement

## Slide 21: What was measured?
Every diode is measured at several (overlapping) positions in the beam

## Slide 22: Scan begin of beam time (decreased current?)

## Slide 23: Scan with more stable current?

## Slide 24: Comparison
In both cases we scanned upwards, meaning each diode's signal on the left was measured first, the one for high y-values (on the x-axis) later on.
The signal is indeed lower for the left scan, which would correspond to the lower beam current. Maybe the beam current dropped after starting the scan and remained nearly constant?
Could we see degradation effects on the right side?

## Slide 25: I calculate a factor for both of these and shift the diodes‘ curves correspondingly

## Slide 26: Factors and comparison
We are, in each case, inside ~ 6 % homogeneity in the response!
There is a difference between the 2 normalization measurements, but some parts of the factor are identical – this is also nice proof of the stability in the diode response.

## Slide 27: Scan Array 3; also stable current – different voltage setpoint
Also, very homogeneous, maybe a bit worse than the other array? Here is one not working diode – the only one of both arrays. This means the fabrication reached an efficiency of ~ 99,6 % 

## Slide 28: Shift back with factor

## Slide 29: Comparison all 3 measurements
We are in a range of ~ 6 % homogeneity for the 2 arrays!

## Slide 30: Now: Compare with older measurement!

## Slide 31
The response is less homogeneous than for the new arrays!
Also: The response is much lower.
From the logbook, these measurements had 1 nA
current, meaning the half of the 2 we had now.
And the setpoint voltages were different:
1.8 now, 2 V back then.

## Slide 32: Comparing the factors
The homogeneity we reached now among one array
is ~ factor 20 times better than before!
Before: Factor from ~ 0.6 to ~ 1.8 -> 1.2 range
120 % deviation in response
Now: ~ 0.96 to ~ 1.02 -> 0.06 range
6 % deviation in response

## Slide 33: Is the new normalisation necessary?
We talk about a signal correction in the percent region…
I add comparisons of maps without (left) and with (right) the new factor correction.

## Slide 34: Round aperture

## Slide 35: Logo

## Slide 36: Misc

## Slide 37: Beam Scan

## Slide 38: Array 3 Logo

## Slide 39: Array 3 Beam Scan

## Slide 40: Maps
Only background subtracted
1st side measurement details + 1 map in real pixels (white space)
2nd side: Maps in pixel (filled)
3rd or 2nd side: Maps in contour

## Slide 41: The round aperture
4 images (1 half out of the beam)
y = 40 mm: 2.43 nA -> 2.34 nA
y = 70 mm: 2.43 nA -> 2.36 nA
y = 85 mm: 2.41 nA -> 2.36 nA
y = 55 mm: 2.43 nA –> 2.38 nA

## Slide 42
y = 40
y = 55
y = 70
y = 85

## Slide 43: The Logo
4 images (over lunch)
y = 55 mm: 2.43 nA ->
y = 65 mm: ->
y = 75 mm: ->
y = 85 mm: –> 2.33 nA
Beam was stable enough to
do automatic measurement

## Slide 45: The misc shape
Shadowing of the effectively smaller ~ 30 mm aperture
4 images (over lunch)
y = 55 mm: 2.43 nA ->
y = 65 mm: ->
y = 75 mm: ->
y = 85 mm: –> 2.33 nA
Beam was stable enough to
do automatic measurement

## Slide 47: The beam shape without diffuser
We actually see a change in the beam between measurement 2 and 3 – this is a prove of our beam imaging ability. Maybe we can make some timelapse of the beam on the next measurement day!

## Slide 48: Beam imaging without diffuser
3 images
y = 55 mm: 2.0 nA -> 1.91 nA
y = 70 mm: 2.0 nA -> 1.95 nA
y = 85 mm: 2.0 nA -> 1.97 nA
Currents at Faraday cup
No diffuser, thus = at sample (if not for the additional collimator in the beamline)

## Slide 50: The Bragg peak (only 1 map, but y translated data)
Map at y = 70 mm
2.44 nA -> 2.44 nA
Y-scan
2.44 nA -> 2.44 nA

## Slide 52: Comparison to previous measurements
with fanout on GaN wafer
without fanout on GaN wafer

## Slide 53: Trapezoid positioning
Latest measurements
Previous measurements
material thickness for max signal: ca. 5.15 mm
material thickness for max signal: ca. 5.3 mm
(Difference will probably arise from signal swap)

## Slide 54: Smaller diodes = higher resolution?
500µm
250µm
200µm
400µm
170µm
SINGLE DIODE
SINGLE DIODE
500µm
200µm
500µm
400µm
400µm
30µm
10µm
10µm
20µm
“Array1“
“Array3“

## Slide 55: Logo Array 3
This was the last measurement of the day, and with the correct set voltage of 1.8 V
Map at y=85
2.45 nA -> 2.42 nA

## Slide 56
Smaller diodes 0.25 mm in scan direction + 0.25 mm steps
0.5 mm diodes + 0.5 mm steps
Looks successful, note that the signal is also nearly exactly halved

## Slide 57: Beam Shape Array 3
Map at y = 85 mm
2.0 nA -> 1.9 nA
(without diffuser)
But with set voltage 1.0 V…

## Slide 58
Quite some change of the beam within our measurement day!
Note that we cannot compare the signal here due to the different set voltages!
Smaller diodes 0.25 mm in scan direction + 0.25 mm steps
0.5 mm diodes + 0.5 mm steps

## Slide 59: Comparison with gafchromic film
Only for diffuser since we haven‘t made a beam image…

## Slide 60: Comparison with diffuser Gafchromic

## Slide 61
H
B
G
D
Looks reasonable, even though we see some curvature of the beam. The grey scale scan
looks more or less flat…
I might try to relate the greyscale and signal values based on scaling.

## Slide 62
H
B
G
D
Norm all to 1 and align x-axes to get a quantitative comparison!
Data from Gafchromic extracted with online tool – points aligned by hand
(without investing several hours this at least gives a nice idea of the data curve)
Looks excellent, and it seems that our normed response can be compared to the response of the gafchromic!

## Slide 63: Ideas for next measurements

## Slide 64: What we have available
New matrix array:
11 x 11 diodes, 2 sizes

## Slide 65: Wedge trapezoid of peak material

## Slide 66: Next measurements
Beam imaging - maybe over a longer duration to generate a sort of “beam video“
More comparisons with Gafchromic! (Several beam shapes, diffuser, maybe also something else? – Bragg, misc, logo?)
New wedge shape + varied proton energy (max signal peak should move on the wedge)
Tests with the matrix arrays: Calibration (moving array through diffused beam), imaging

## Slide 67: Standard for file naming
I think it might be a good idea to have a standard for the file naming. Currently, the naming is good – but some slight problems still exist.
My suggestion:
{#Number of measurement}_{Short name of measurement / object}_{Diffuser_thickness}_um_{Current at beam level!}_nA_{Set voltage of ams circuit in V}_x_{x pos of stage}_y_{y pos of stage}.csv
e.g.: 1_RoundAperture_200_um_2,43_nA_1,8_x_20_y_70.csv
Thus each measurement is distinct (due to the numbering) and we keep the most important values (easy to parse) in the file name.
Adapting this should be easy with Stéphanes Python script.