## Slide 1: Results Cyrcè 19.06.24

## Slide 2: Changes in data treatment
All diodes are recorded – no excluded channels
  Compared to the last measurements the channel numbering is mirrored?!
File positions are given different (clearer) as _x_..._y_....csv
  new parsing of position was written
No variation in beam: Otsu‘s method for generating an on-signal and dark measurement is not available any longer
One ?! dark measurement: “d2_1n_3s_beam_all_without_diffuser_dark.csv“
  General usage of dark measurement added
Normalization factor: Compare following slides
Channel 60 with extremely high noise (but no proton signal) –> reduced or not used in following

## Slide 3: Dark measurement
Dark measurement
Applied on higher signal

## Slide 4: Normalization – what‘s measured?
Let‘s first look at what measurements are available for the calibration.
We have two measurements with varied y-position and constant x-position. These can be used effectively for the calibration since we can consider many different points with beam signal for each diode.
I will show an overview in the next slide – just for a quick impression – of how much the diode array moved in these measurements. (The crosses you see mark the measurement y-position; for each point I sketch the diode array in original y-dimension. These sketches are shifted on the x-axis.)
Why are there 2 graphs? We have to suited measurements, but one (the right one, earlier at the day) was assigned with an unstable beam and might thus not be a good choice for the calibration. Nevertheless, I‘ll show both measurements.

## Slide 5: Normalization – what‘s measured?
Available: 1 measurement at beginning and one at the ~ end of the measurement day – unstable beam for the first measurements on the 19.06.24
500p_center = begin of measurement day
5s_flat_calib = end of measurement day

## Slide 6: Normalization
If we analyze the measurements, a problem occurs. The y-position recorded for one measurement is the same for all diodes.
What we need for the normalization: Compare the signal from diodes if they are at the same “real“ position.
If we do not consider the real position the following problem occurs:
(About the plot format: I will show the signal amplitude vs. the y-position for all diodes at once. To decide between the diodes I assign a colourscale to the diodes numbered #1 to #64)

## Slide 7: Signal vs. y-position
Result:
With these information it seems hardly possible to find a calibration…

## Slide 8: Normalization
If the beam was flat, we would not have any further problems; we‘d just take the diode signal at one position in the beam or mean all positions in the beam.
BUT: The beam is not flat, we observe a amplitude dependence for the diodes‘ curves with the y-position.
To find a meaningful calibration it is then necessary to compare the diodes‘ signal at the same position of the beam.
This is possible with two approaches:
  We try to find the real diode position from the signal (all diodes should align in their beam on / beam off edges. (Might cause problems for diodes with low signal)
  We use the position information we have about the diodes, fabricated with a clear defined optical mask. (Favored!)

## Slide 9: Signal vs. real y-position
Result:
The flattened beam behind the diffuser is not flat at all (shape comparable over the day)
The diodes show a differing response to the same signal -> Use this for normalizing

## Slide 10: Normalization
As visible this works – we observe a nearly uniform shape of the beam profile for the different diodes!
The signal of the diodes shows a different height – this is what we want to cover with the normalization!
If the recorded curve shape differed between the diodes, this would indicate a not-uniform signal response of the diodes. (We do not expect this, because the measured linearity of this array was nearly 1!)
Next, I summarize the steps of the normalization. Note one challenge for this: Not all diodes cover the whole beam (thus some curves begin and end in the middle of the beam). We need to find some possibility to compare all diodes at the positions they‘re recorded.

## Slide 11: How to realize new normalization
Calculate the real position for each diode at each y-position
Mean the signal at each real position
  To do this: Group positions in a certain range
For each diode: Calculate how the measured beam curve differs from the mean
I allow the following DOF : mean = signal diode * a
  Realisation with different methods possible, currently I use a least_squares optimisation
  More DOF are possible, then we‘re not only dependent from the channel, but also from the y-position (Assuming the linearity is - as measured before - nearly , there shouldn‘t be a big impact though)
Thus a factor is calculated for each channel

## Slide 12: Some technical aspects
The factor is limited (to avoid artefacts in channels with no response during flat calibration): For the results shown to 0 and 3
The factor can be normalized, so that mean(factor) = 1 (Also done for results shown)
This calculation process is more complex (~ 5s), thus i built a quick storage as “normalization_factor.npy“ in the given data folder via np.save/np.load (< 0.001 s)
  I always safe a raw factor without the 2 modifiactions above, to allow later modifications of these after loading
This process is rather complex, but I could not think of a more direct way ... So maybe you have some ideas/questions

## Slide 13: The real positions
Here a visualisation how the real y-positions are calculated and grouped (I use the distance of 1xdiode size to group)
The signal increase is a bit smeared out (compare next slide), so there could be some offset in the diode array? (how accurate is the fabrication process) or some inaccuracy in moving the array (e.g. if the array is mounted slightly tilted)
Future Improvement for this: Detect the signal increase/decrease and rather use this for sorting

## Slide 14: The new normalization factor
In black the calculated mean, As visible I‘m able to get the diode signals far more homogenous

## Slide 15: Normalization with normalized factor
With mean(factor) = 1 we of course get an offset from the mean here

## Slide 16: Factor comparison
Comparison of the factors for 2 methods(lsq = least_squares, method 1 = mean of deviation at each position compared to mean)
Also the factor from the first Cyrcè measurements is shown in green
The methods give nearly the same result
The behaviour between the two measurements at the 19.06. is differing, but especially for the diodes with #30+ not too much different

## Slide 17: One addition that might be important?!
In the recorded images: Correct for the beam shape in y-direction (but this assumes that it remains comparable over the day)
(If it‘s not comparable the normalization is of course still valid, since it just covers the difference in the diodes‘ response. This won‘t change with other beam shapes.)
Possible if assigning the measured real position here to the y-positions in the maps measured

## Slide 18: The maps
I will show the maps with/without normalization in the following with a short picture context to each measurement
Measurements are shown in realistic pixels and interpolated contourf version (as before)
If multiple y positions were measured, I‘ll show 2 maps
  Combining the two will be one of the next steps
For the first few maps, I will show on the left side a normalized and on the right side a raw data image – to judge if the normalization improves the images.

## Slide 19: IPHC Matrix CRHEA logo

## Slide 20: IPHC Matrix CRHEA logo
Normalized
Not normalized

## Slide 21: IPHC Matrix CRHEA logo 2 (“neonoff“)
Normalized

## Slide 22: IPHC Matrix CRHEA logo

## Slide 23: IPHC Matrix CRHEA logo 2 (“neonoff“)

## Slide 24: Mixed symbols mask

## Slide 25: Mixed symbols mask
Here the normalization seems to be an improvement

## Slide 26: Mixed symbols mask
Here the Bragg-peak effect is well visible I‘d say

## Slide 27: Biseau 2D-big black trapeze

## Slide 28: Biseau 2D-big black trapeze

## Slide 29: Biseau 2D-big black trapeze

## Slide 30: white trapeze – biseau blanc
picture only with rat tails

## Slide 31: white trapeze – biseau blanc

## Slide 32: white trapeze – biseau blanc
picture only with rat tails

## Slide 33
From now just normalized maps, left/right for height or picture comparison

## Slide 34: Rat tails on the white trapeze – biseau blanc

## Slide 35: Rat tails on the white trapeze – biseau blanc

## Slide 36: Rat tails on the white trapeze – biseau blanc

## Slide 37: Rat tails on the white trapeze – biseau blanc

## Slide 38: Rat tails on flat surface

## Slide 39: Rat tails on flat surface

## Slide 40: Rat tails on flat surface

## Slide 41: Beam without diffuser

## Slide 42: Beam without diffuser

## Slide 43: “Full scan“ 500pA with diffuser (and unstable beam?...)

## Slide 44: Some remarks
I think there might be a differing beam shape behind the diffuser in the x-direction as in the y-direction
  From the maps and beam_shape image I‘d say less intensity on the left side of images
From what I see on the pictures the objects seem to be resolved as expected

## Slide 45: Some ideas for future measurements (if there are anymore with a 1D array)
More measurements at slightly shifted y-positions (e.g. to get a better result for the IPHC MATRIC CRHEA Logo)
Maybe an improved logo (clearer cutout in e.g. “I X“, shift towards Bragg peak for more signal?)
Curved shapes still strange resolved: Examine the impact of diode movement
  Map something 2D so that a map out of one diode can be generated
  Is this showing the same effect now also in the y direction?
  (Also possible with rotated diode array)

## Slide 46: What‘s coming in the next days
Black trapeze to estimate Bragg peak (if i get the measures)
Change in software from script function approach probably into a class based approach for better handling (and easier modifications)
Software for 2D array measurments (maybe with interactive features so that its usable when we do the next measurements)