## Slide 1: Results 11/10/24 Results 22+23/10/24
I will mix between these measurement days, since I will focus on the different topics in the analysis.

## Slide 2: What was measured?
11/10/24: Array 128x0.5x0.5 mm2
y-translation through undiffused beam for normalisation under difficult circumstances?
Multiple pictures of undiffused beam, and the 3 diffusers with Gafchromic scans afterwards for comparison
New Bragg wedge out of PEEK material – combined with different energies from Al foil wheel
9 scans of undiffused beam over 1 hour
After second scan of beam there was a crash of the array into the metal parts of the proton outlet (likely to due tilt of stage from touched stage connction cables). We saw a saturated current, luckily there was NO! fatal damage to any diode or to the readout circuit. If some properties of either of those changed is of course not directly clear and will be investigated.

## Slide 3
11/10/24: matrix 11x11 1x1 mm2 #1
Voltage scans
x and y-translation to diffused beam
xy scan of beam to puzzle into one picture (with less time requirement than scan of linear array)
movie scan of beam (with fixed position, but several changes of the beam during the one long scan ~1500 frames over ~ 20 minutes)

## Slide 4
21/10/24: matrix 11x11 0.5x0.5 mm2 #1
Voltage scans + voltage dependent linearity
y-translated scan
xy scans of round aperture and star of misc shape with steps of half the diodes size for super-resolution
movie scan of beam over lunch (unchanged beam) + Gafchromic
scans of middle of undiffused beam at different proton energies from wheel (7 energies from 23.7 to 3.5 MeV) – what changes in terms of signal to noise ratio? What can we learn about our energy resolution?
y-translation after all these measurements
linearity afterwards
image of the misc star at the 7 chosen wheel positions – image quality in energy dependence
PEEK wedge for all 20 wheel positions with y-scan
21/10/24: array 2x64 0.5x0.5 mm2 #1
PEEK wedge for all 20 wheel positions (stationary)

## Slide 5
22/10/24: array 2x64 0.5x0.5 mm2 #1
voltage scans
y-translation for normalization
linearity in voltage dependence
Gafchromic comparisons of: all three diffusers (40, 200, 400 um); logo, misc, PEEK wedge, mouse-foot, undiffused beam
22/10/24: matrix 11x11 1x1 mm2 #2
voltage scans
y-translation for normalization
movie of beam over lunch
movie with increased current (increase each 30 seconds, ~1fps, from 100 pA to 2 nA)
linearity test, voltage dependent
another 2 movies of undiffused beam with some live changed features
movie of the moved logo (sadly with some problems due to improvised change of the stage translation – more details later – but at least we filmed the crash of the stage)

## Slide 6: Strange 50 Hz modulation only with diffuser

## Slide 7: Some important remark
During the measurements with lower integration time with higher frame rate I observed some periodic pattern on the signal with samples) meaning on the time axis!
I observed this more in detail, and will continue to make comparisons for all measurement days.

## Slide 8: 50 Hz noise on undiffused beam?!
Compared is the highest and lowest (dark) signal of arbitrary measurements with and without diffuser
I observed this modulation for different measurement days whenever we measured without diffuser; I will nevertheless bring these observations in plots for you to compare
If you count the periods on the right side you find: 10 periods in 200 ms meaning 50 in 1s = 50 Hz…

## Slide 9: Where could this come from?
We first thought about the readout electronics, but this is not explaining why we observe only without diffuser
The Cyrcé guys thought about a pump and a mechanical vibration – it was not the cause and would also not explain the diffuser behaviour
Obviously the fact that the modulation disappears when the diffuser is in the beam is really odd
(I started the following thought: A diffuser should compensate for the time modulation of a proton beam because it changes the trajectories of incoming protons. You can continue the thought and if this would imply something for the undiffused proton beam at Cyrcé…)

## Slide 10: New voltage scan results

## Slide 11: Some remarks
I made some mistake with the results I have shown earlier for the 128 diode line arrays!
The problem was: I considered all diodes, but we only had a part of these arrays in the beam
Thus, I wrote a new routine for voltage scans (since this is part of our standard characterisation of arrays) with implemented filter for signal diodes (this also excludes not properly working diodes)

## Slide 12: Voltage scan up to the 11th October

## Slide 13: Sensitivity and signal quality

## Slide 14: And most importantly the ratios!
These are the ratios normed to the smallest array (128x0.25x0.5)
A dashed line shows the caluclated ratio when considering the active region of the diodes (meaning 0.4 x 0.4 mm2 for the 0.5 x 0.5mm2 diodes)
These new results are remarkable and different from what I presented earlier. The new implication is: The detection seems to happen exactly in the region we intend it to happen!
For the matrix array, we observe a discrepancy from the estimated ratio, but I will explain later why this is not really comparable.

## Slide 15: All voltage scans up to the 22nd of October

## Slide 16: Sensitivity and signal quality
Note that the sensitivity implies that we could e.g. detect a current ~ 200 lower than we had for the big matrix arrays)
Meaning: 2 nA at Faraday cup / 200 = 10 pA at Faraday cup under these conditions for the large diodes.

## Slide 17: All ratios
All arrays normed to small 128x0.25x0.5 array, but conditions are not all comparable!
All new arrays from 21/22 = comparable, normed to small matrix.

## Slide 18: Still not really correct
I wrote in the header of the plots @ 2 nA, but this is false
On the 23rd of September, we had 2.43 nA at Faraday cup to get 2 on the sample and the additional collimator
Afterwards, we kept 2 nA at the Faraday cup…
To make really correct plots, I need to correspond to the proton current we really had on each measurement day. So far, the info we got from the Cyrcé crew does not really fit, so I am waiting for further updates.

## Slide 19: Movie time: logo and current increase
Speed by factor 20 increased to 20 fps.
If movies are not displayed correctly in PPT: I put them as .zip into the results folder

## Slide 20: Movie current increase
Needs refinement, like normalisation and eventually filtering for uniform intensity scale
We used a square aperture shadowing the borders of the array.
The array is slightly tilted.

## Slide 21: Movie Logo move
The following happened around/ during the movie:
Planned was an x-translation (photo) but this was not possible due to the extent of the linear stage over the vacuum box
The change to y-worked mechanically, but the code was not ready for it. Somehow the controller has a factor 2 on this y-axis stage, which was not considered.
From design the movie is reversed: Meaning we started with the end of the logo and ended the movement at the intended start of the movie.
Thus we started the linear movement far too late and moved far too wide with the y-stage.
We only have the IPHC fully in the movie, CRHEA is cut on the first letter (I will not interfere in this Strasbourg – Nice thing here )
But the start of the movie shows (for a relatively long time) how the stage moved into the 2 limits and broke them… - live documented!

## Slide 22
I still need to compensate for the not working pixels, normalize everything, and mirror the image (since it‘s false from this perspective).
Some other remark: The logo was too thin and thus did not absorb the proton beam. We thus have a contrast from red (protons with lower energy) to grey (open proton beam). And some density fluctuations in the logo make parts harder to recognize.

## Slide 23: General remarks

## Slide 24: General stuff
Logbooks will be scanned and made available for you.
We need to discuss conferences/papers, as we already talked about with Matilde, Maxime & Stéphane.

## Slide 25: Some plot-style remarks (from your side?)
I can quickly change the colour scale, format, etc. , of my result plots (made with Matplotlib). So if you have some remarks or suggestions, … please let me know. Maybe we can discuss some uniform plot styles we use in the MATRIX project.
E.g.: I can put every colour scale for the intensity of our images/videos. I suggest some perceptually uniform scales as described by https://seaborn.pydata.org/tutorial/color_palettes.html
(I try to avoid the quite often used rainbow scale because of the green part that's recognised more by the eye.)
It is also possible to design any colour scale on our own by just handing Matplotlib the colours it should vary in between.
Some examples: These two are uniform and avoid light or dark shades that could vanish before a background.

## Slide 26: Linearity
Standard process for analyzing
Results small and big matrix arrays 21+22/10/24
Is the voltage chosen impacting the linearity of our arrays?

## Slide 27: What will be analyzed
We measured the linearity voltage-dependent for currents between 0.5 nA – 2.5 nA (5 points)
This is not the biggest range, but it is enough for some quick statements about the linearity of arrays (especially when collecting measurements as a benchmark for new arrays)
Again only diodes with signal will be considered! Signal = netto signal (signal –dark)
I will plot the mean signal of all diodes (black, x) and the signal of each diode (grey, x)
Also plotted: Dark noise (black, v); mean of signal std (blue, ^); std of diodes signal (green, ^)
I also show the inaccuracy of the mean signal = Sqrt[mean of signal std2 + std of diodes‘ signal2] (black, ^)
I fit a linear function of type f(x) = a*x to each diode's response and collect their R2 value – in the plot I will show the mean of these fits (red, -- line) and the mean of all R2 values.
I also fit a function of type f(x) = a*x1/2 + (dark noise) to each signal noise, again I will show the mean (red, -- line) and the mean of all R2 values.

## Slide 28: Small matrix # 1
From 0 until 0.6 V the arrays show no meaningful response as the current detected is saturated

## Slide 29
From 0.7 V we detect a real signal (as seen in voltage analysis); from 0.9 V the detected signal behaves as wanted with excellent linearity

## Slide 31
For 2 V the linearity is still excellent, but the noise becomes so small that it is not always detected. (We observe many 0‘s in the file) Thus, the signal std is not any longer growing with sqrt of the signal as we‘d expect.
This, together with the voltage analysis, serves as justification for the voltage choice of 1.9 V for nearly all measurements we did at Cyrcé since September!

## Slide 32: Comparison between voltages
Shown in blue is the agreement between the expected model and the real behaviour of the diode signal (signal noise).
For all voltages from 0.9 V the linearity is excellent (R2 = 0.99x). There is not a real tendency observable with changing voltage.
The agreement with the sqrt growth of the signal noise is less good, but still in a decent range 0.95 – 0.99.
Shown in red is the spread of the diode‘s signal (as std of diode response/mean of diode response for each current point, and averaged)
For most voltages, this is around 9 %. (Note that the diodes are at different points in the diffused, but not totally flat, beam).

## Slide 33: Big matrix # 2 (measurements started at 1 V)

## Slide 35: Comparison between voltages
We again observe excellent linearity for all voltages from 0.9 V (R2 = 0.99x) , possibly with a slight increase for higher voltages.
The agreement with the sqrt growth of the signal noise is again less good, still in the range 0.95 – 0.99. (again except for 2 V)
The inhomogeneity in the diode‘s response is a bit higher at ~ 13 %, with a slight increase towards higher voltages. (We have bigger diodes (twice the active region) and still the diodes‘ signal might differ due to the characteristics of the diffuser.

## Slide 36: Résumé
The linearity is excellent (for the small current range tested)
The signal noise is also behaving as expected
The linearity is excellent for all voltages from 0.9 V, maybe with a slight improvement towards bigger voltages.
2 V is limited in the recording of the (signal) noise
The inhomogeneity in the diodes‘ response shows no dramatic change with the voltage chosen. Except for 2 V being clearly worse than 1.9 V. (Smaller voltages could be a bit better regarding this aspect?)

## Slide 37: Images with the 2 x 64 0.5 x 0.5 mm2 array

## Slide 38: Mapping
From the images of the array I have, I assume to be the bonding as shown:
Additionally, there will be the mapping between bonding / PCB and readout circuit.
The right image is the top view on the array, meaning in the measurements this is from the beam perspective. In the measurement, we moved the array from the beam
perspective to the left (from our view to the right)– thus, the left line of the array will be aquivalent to the line further into the moving direction.
For image generation, I will thus sort the diodes into two 64-long arrays with an offset of 0.25 mm (this option needs to be added to my Analyzer class). Because we look at the images from our view – opposite to the beam direction – the right array will contain the diodes further into a moving direction (left in the right schematic)
I will first generate two images with each line separate, then use the super-resolution available. For this, I will consider the signal of the diodes at the same position and double the number of pixels in the y-direction.
1
2
3
4
…

## Slide 39: Results without super resolution
1 working but too high and one dead diode – could be from readout circuit
The resolution of the mouse foot is excellent!

## Slide 40
2 images 0.25 mm in y-shifted for super-resolution

## Slide 41: More images + superresolution coming
1 Problem we have for this array: The array was tilted, thus all measurements show this slight angle.